## Specifications
Product | Valiant 2 rehab XXL
Partnumber | 938915
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Maximum speed | Maximum speed of the treadmill | 20 km/h | 12.4 mph | x
| Minimum operational speed | Minimum speed of the treadmill | 0.1 km/h | 0.1 mph | x
| Positive elevation | Maximum positive angle as a percentage | 15 % |  | x
| Elevation adjustment steps | The size of the steps you can adjust the elevation | 0.5 % |  | x
Accuracy | Speed accuracy | Accuracy of speed | 5 % |  | x
| Accuracy inclination | Accuray of the inclination | 0,5 % |  | x
Comfort | Allowed user weight | Maximum patient weight | 225 kg | 496.0 lbs | x
User Interface | English user interface | User interface available in English language | yes |  | x
| Norwegian user interface | User interface available in Norwegian language | yes |  | x
| Czech user interface | User interface available in Czech language | yes |  | x
| Danish user interface | User interface available in Danishlanguage | yes |  | x
| Dutch user interface | User interface available in Dutch language | yes |  | x
| Finnish user interface | User interface available in Finnish language | yes |  | x
| French user interface | User interface available in French language | yes |  | x
| German user interface | User interface available in German language | yes |  | x
| Italian user interface | User interface available in Italian language | yes |  | x
| Japanese user interface | User interface available in language | yes |  | x
| Korean user interface | User interface available in Korean language | yes |  | x
| Polish user interface | User interface available in Polish language | yes |  | x
| Portugese user interface | User interface available in Portugese language | yes |  | x
| Russian user interface | User interface available in Russian language | yes |  | x
| Spanish user interface | User interface available in Spanish language | yes |  | x
| Turkish user interface | User interface available in Turkish language | yes |  | x
| Ukrainian user interface | User interface available in Ukrainian language | yes |  | x
Connectivity | Optional USB connector | Connection via USB possible | yes |  | x
| Optional RS232 connector | Optional connection via RS232 | yes |  | x
Dimensions | Screen resolution | Screen resolution | 800 x 400 pixels |  | x
| Walking surface length | Length of the walking surface | 250 cm | 98.4 inch | x
| Walking surface width | Width of the walking surface | 120 cm | 47.2 inch | x
| Step up height | Height to step up the treadmill | 17 cm | 6.7 inch | x
| Product length (cm) | Length of product in cm | 312 cm | 122.8 inch | x
| Product width (cm) | Width of the product in cm | 148 cm | 58.3 inch | x
| Product height | Height of product | 130 cm | 51.2 inch | x
Power requirements | Power cord length | Length of powercord including plugs | 250 cm | 98.4 inch | x
| Power cord IEC 60320 C19 with CEE 7/7 plug | European power cord and connector standard supplied with product | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
| Maximum rated current | Maximum rated current | 12 A |  | x
| 115 V AC 50/60 Hz (2 phases) | When <230V, the treadmill should be connected between two phases | yes |  | x
| 230 V AC 50/60 Hz | Operational power | yes |  | x
| Maximum motor power | The maximum power the motor will consume | 3.0 kW |  | x
Standards & Safety | IEC 60601-1:2005 | The product is IEC 60601-1 3rd edition compliant | yes |  | x
| ISO 13485:2016 compliant | Lode fulfils ISO 13485:2016 requirements | yes |  | x
| ISO 9001:2015 compliant | Lode fulfills ISO 9001: 2015 requirements | yes |  | x
| Standard emergency lanyard | An emergency stop with a magnetic lanyard is present | yes |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 40 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 106 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 70 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | 70 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | -25 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 95 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 10 % |  | x

## Required options
Name |  | Partnumber | Required
Remote Control for treadmill | Ease of use | 945831 | x
Add program function to 7" touch screen for Treadmill | Programmable | U945815 | x
Remote Control for treadmill (ordered afterwards) | Ease of use | P945831 | x
Table foot for Control Unit with Touchscreen | Control from any desired place | 945818 | x
Cable for separate Control Unit Treadmills | Multifunctionality | 945813 | x
Colour Display 3.5" - 2nd screen | Multifunctionality | 945819 | x
SpO2 for control unit with touch panel (treadmill) | Oxygen saturation | 945822 | x
Adult Finger clip, 1 meter cable | Easy to apply and clean | 928946 | x
Adult Finger clip, 2 meter cable | Easy to apply and clean | 928947 | x
Adult Finger clip, 3 meter cable | Easy to apply and clean | 928948 | x
Earclip sensor, 1 meter cable | Easy to apply and clean | 928951 | x
Reflectance sensor, 1 meter cable | Easy to apply and clean | 928952 | x
Reflectance sensor holder, 10 caps & 20 adhesive stickers | Easy to apply and clean | 928953 | x
Adult flex sensor with 25 flexiwrap, 1 meter cable | Easy to apply and clean | 928954 | x
Adult flex sensor with 25 flexiwrap, 3 meter cable | Easy to apply and clean | 928955 | x
Sensor Adapter Cable, 18 cm length, 8 to 9 pin connector | Versatile connectivity | 928961 | x
Reusable Soft Sensor Medium, 1 meter cable | New technology | 928967 | x
Reusable Soft Sensor Medium, 3 meter cable | New technology | 928969 | x
SpO2 for control unit with touch panel - ordered afterwards | Ordered afterwards | P945822 | x
Blood Pressure Measurement with ECG trigger for treadmills | with ECG trigger | 945824 | x
Blood Pressure Custom pack 01 | For GE Case - GE Cardiosoft ECG CAM 14 - Case 8000 - Sensormedics Max 1 - Mac 5000 | 945881 | x
Blood Pressure Custom pack 02 | For Burdick Quest - Cambridge Heart CH2000 (V1.5 to 2.0.3) - Cosmed Quark T/C12x - Marquette Case 12, 15, 16 / Centra - Mortara X-Scribe II - Norav - Philips Stress Vue | 945882 | x
BPM Custom pack 03 | For Marquette MAC VU Digital ECG | 945883 | x
BPM Custom pack 04 | For Marquette Hellige Cardiosys | 945884 | x
BPM Custom pack 05 | For Quinton Digital ECG - Nasiff Associates Cardio-Card | 945885 | x
BPM Custom pack 06 | For Welch Allyn CardioPerfect Workstation | 945886 | x
BPM Custom pack 07 | For Reynolds CardioDirect 12-S | 945887 | x
BPM Custom pack 09 | For Schiller CS200 | 945889 | x
Cuff Complete Adult - small (18-27 cm) | easy to apply | 945870 | x
Cuff Complete Adult (25-35 cm) | easy to apply | 945871 | x
Cuff Complete Adult - plus (27-40 cm) | easy to apply | 945872 | x
Cuff Complete Adult - large (32-44 cm) | easy to apply | 945873 | x
Emergency Stop Button | Ultimate safety | 945804 | x
Entrance plate for Valiant rehab XXL | Even easier entrance to the treadmill | 938849 | x
Ambient sensor pack | Check environmental conditions during test | 945826 | x
Extension for emergency lanyard | Original Accessory | 945931 | x
USB to Serial converter | Easy connection | 226012 | x
Network module connection cable | Easy connection | 930930 | x
Communication Module | Connect to ECG and spirometers | 945850 | x
RS232 cable | Easy connection | 930911 | x
Handrails, Side - Adj. for extended Valiant 2 | Safe and flexible handrails | 938821 | x
Bluetooth Smart heart rate | Available in Q2 2019 | 945833 | x

## Spareparts
Name | Partnumber
3,5" Control Unit front panel | 945110
7" Control Unit front panel | F945105
7" Programmable Control Unit front panel | F945105P
Frequency controller | 250002
Inclination sensor | 9457433
Main switch with fuse | 217201
Treadmill Control Module 2 (TCM2) | F945705